
[image: image1.png]


412                     

EKEM2D
Installation and operation instruction T1.1

  ZHEJIANG ETEK ELECTRICAL TECHNOLOGY CO.,LTD.
Declaration
The copyright is the property of ETEK. Any information in any paragraph or section cannot be extracted, copied or otherwise reproduced or propagated. Otherwise offenders shall take all consequences.

All rights are reserved.

ETEK reserves the right to modify the product specifications herein without notification. Please consult the local agent about the latest specifications before placing a purchase order.
Contents 
- 1 -

1 General

- 1 -

2 Function

- 1 -

3 Technical parameter

- 3 -

4 Outline (unit: mm)

- 3 -

5 Wiring and installing

- 3 -

6 Analysis and elimination of common faults

- 4 -

7 Operation and display

- 8 -

8 Communication description



1 Overview
EKEM2D single phase electric meter is designed for single phase active energy measurement on low voltage system, in the same time it can measure the electrical parameters like voltage, current, power and so on. There is also RS485 can be chosen. This electricity meter has advantages of smaller volume, high precision, good EMC, easily installing etc, All meters meet the related technical requirements of electricity meter in the IEC62053-21、IEC62053-22 standards.
2 Function
	Function
	Function description
	Function provide

	Measurement of kWh
	Single-phase active kWh (positive and negative)

	■

	Measurement of electrical parameters
	Voltage, Current, Active power, Reactive power, Apparent power, Power factor and Frequency
	■

	LCD Display
	8 bits section LCD display
	■

	Key programming
	3 keys to set parameters like code, address, baud rate, multi-tariff and communication protocol 
	■

	Pulse output
	Active energy pulse output
	■

	Multi-tariff
	Date and Time

3 months historical energy data frozen storage
Adapt 4 time zones, 2 time interval lists, 14 time interval by day and 4 tariff rates
	□F

	Communication
	Communication interface: RS485, Communication protocol: MODBUS-RTU
	■


(■：Standard；□：Optional)

3 Technical parameter
3.1 Electric performance
	Input voltage
	Reference voltage
	AC 230V

	
	Reference frequency
	50Hz

	
	Power consumption
	<10VA

	Input current
	Reference current
	0.5-10(80)A

	
	Consumption
	<4VA

	Measurement performance
	Accuracy of measuring
	Class B

	
	Range of measuring
	000000.00~99999999kWh  

	 Clock accuracy
	Error≤0.5s/d 

	Active pulse
	Pulse width
	80±20ms

	
	Pulse constant
	1000imp/kWh


	Communication
	Interface
	RS485(A+、B-)

	
	Connection mode
	Shielded twisted pair conductors

	
	Protocol
	MODBUS-RTU


3.2 Mechanical performance
	Outline 
	Length × Width × Height
	90mm×36mm×65mm 

	Strong current terminal

Torque
	<1.8Nm


3. 3 Work environment
	Temperature range
	Work temperature
	-25℃ to +55℃

	
	Storage Temperature
	-40℃ to +70℃

	Relative humidity
	≤95%(No condensation)

	 Altitude
	<2000m


4  Outline (unit: mm)
[image: image25.png]



Meter outlook and size
5  wiring and installing

[image: image2.png]-

L

L

N N

21 2217 18

A B + -
RS485

Load

Active

Pulse




6  Diagnosis, analysis and elimination of common faults
6.1 Auxiliary power failure

Failure performance: the meter flashes and does not light up after being powered on.

Troubleshooting: 1. Check whether the wiring of the auxiliary power supply is consistent with the wiring diagram of the instrument, and whether the wiring is loose or falling off;

2. Use a multimeter to measure whether the input voltage value of the auxiliary power  supply is within the normal working voltage range of the instrument.

6.2 Signal input failure
Failure performance: After the meter is powered on, the display power or energy count is not accurate.

Troubleshooting: Switch the display interface of the meter to the power (active P, power factor λ) interface, check whether the power display is negative and whether the power factor is between 0.9-0.95, and then check whether the input and output of the current signal line are reversed ( That is, the incoming line of the current must be consistent with the incoming end of the instrument), And consistent with the wiring on the meter.
6.3 communication failure
Failure performance: After the meter is powered on, it cannot communicate with the host computer normally.

Troubleshooting: 1. The voltage value between the communication output A and B of the measuring instrument should be between +(4.4-4.5)V;



2. Check whether the communication wiring method is correctly wired according to the wiring diagram (that is, the communication terminal A/B of the instrument should correspond to the communication serial port A/B);

7  Operation and display
7.1 Key description
	Key icon
	Key name
	Key　function

	 
[image: image3.png]



	Key up
	View voltage and current in the view interface
Up and flashing shift in the programming interface

	 
[image: image4.png]



	Key down
	View power in the view interface
Scroll down and modify flashing bits in the programming interface

	
[image: image5.png]



	Key setting
	View electrical energy in the viewing interface

Long press 3S to enter/exit the menu

Short press OK in the programming interface to save the settings


7.2  display description
Show total energy when connected. Change information while pressing down key. Display information as following:
	
[image: image6.png]



	U、I、F、Time、MODBUS　Address、Baud、parity、Version、ALL－display；

	
[image: image7.png]



	Total　active　power、total　reactive　power、total　apparent　power、total　power　factor；

	
[image: image8.png]



	Total active energy, forward active total energy, reverse active total energy, total active spike energy, total active peak energy, total active flat energy, total active valley energy, total reactive energy, forward reactive total energy, reverse total reactive energy, total reactive spike energy, total reactive peak energy, total reactive flat energy, total reactive valley energy.


Note：

Listed above are the names of all display interfaces of the EKEM2D meter with double rate function. Three buttons can switch different types of display content, the switching sequence is as described above；

For the EKEM2D meter without the double rate function, it does not display the date, time and various types of time-sharing energy (the energy in the four rate periods of sharp, peak, flat and valley).
	[image: image9.png]



	[image: image10.png]




	Current total active energy(6961.96kWh)
	Current total reactive energy

	[image: image11.png]



	[image: image12.png]




	Voltage
	Current

	[image: image13.png]



	[image: image14.png]




	Current total active power 550W 
	Current total reactive power 953VAr

	[image: image15.png]



	[image: image16.png]L
J






	Current total apparent power 1.100kVA
	Current total power factor 0.500


Note:The above is just a part of the display interface. The display mode of other interfaces is similar to the above figure. You can judge the display meaning according to the information displayed on the interface.
7.3Programming display menu
Press 
[image: image17.png]


 at any main menu and get in [image: image18.png]


 interface, and then press 
[image: image19.png]


 show[image: image20.png]


, and enter the code. If you enter a wrong code, it will show “0000” and enter the code again; and if you enter a right code, you can set the parameter. After setting the parameter, it will show [image: image21.png]


 and save the change by pressing [image: image22.png]L™




 and quit without save by pressing[image: image23.png]2
ng




.
[image: image24.png]l
~Addr O .bAL wPRE

i 88 nanf
I TR |

dE Press (2] Enter the pogeto beset
n
u

f:::; mSHUE E MS




7.4 Item can be set 

Setting item description
	Mum
	Secondly menu

	
	Symbol
	Meaning
	Range

	1
	ADDR
	Communication address
	1-247

	2
	Baud
	Baud setting
	1200、2400、4800、9600、19200、38400

	3
	Pari
	Parity setting
	None、Odd、Even

	4
	LED
	Background light setting
	0-255 minutes， 0 ever bright

	5
	S-TY
	Apparent power calculation
	PQS,RMS

	6
	EF-E
	Set multi-tariff
	EF-YES

E-NO

	7
	DIR
	Current direction
	no-forward

 yes-reverse

	8
	CoDE
	Code setting
	1-9999


8 Communication description
8.1 Communication protocol
The meters adapt Modbus. Please refer to the relevant standards for more information. The multi-tariff data mean nothing when multi-tariff function (F) is not applied.
8.2 MODBUS Address list
	Address
	Variable
	Length
	Attributes
	Note

	0000H
	Current combined total active energy
	4
	R
	unit：0.01kWh

	0002H
	Current combined spike active energy
	4
	R
	

	0004H
	Current combined peak active energy
	4
	R
	

	0006H
	Current combined flat active energy
	4
	R
	

	0008H
	Current combined valley active energy
	4
	R
	

	000AH
	Code
	2
	R
	

	000BH
	Voltage
	2
	R
	unit：0.1V

	000CH
	Current
	2
	R
	unit：0.01A

	000DH
	Active power
	2
	R
	unit：0.001kW

	000EH
	Reactive power
	2
	R
	unit：0.001kvar

	000FH
	Apparent power
	2
	R
	unit：0.001kVA

	0010H
	power factor
	2
	R
	unit：0.001

	0011H
	Frequency
	2
	R
	unit：0.01Hz

	0012H
	Year, month
	2
	R/W
	

	0013H
	Day, hour
	2
	R/W
	

	0014H
	Minute, second
	2
	R/W
	

	0015H
	Address
	1
	R/W
	1~247

	0015H
	Communication baud rate
	1
	R/W
	00:1200

01:2400

02:4800

03:9600

04:19200

05:38400

	0016H
	light time
	2
	R/W
	

	0017H~0021H
	Reserve

	0022H
	Total active energy of last month
	4
	R
	unit：0.01kWh

	0024H
	Spike active energy of last month
	4
	R
	

	0026H
	Peak active energy of last month 
	4
	R
	

	0028H
	Flat active energy of last month
	4
	R
	

	002AH
	Valley active energy of last month
	4
	R
	

	002CH
	Total active energy of last 2 month
	4
	R
	

	002EH
	Spike active energy of last 2 month
	4
	R
	

	0030H
	Peak active energy of last 2 month
	4
	R
	

	0032H
	Flat active energy of last 2 month
	4
	R
	

	0034H
	Valley active energy of last 2 month
	4
	R
	

	0036H
	Total active energy of last 3 month
	4
	R
	

	0038H
	Spike active energy of last 3 month
	4
	R
	

	003AH
	
Peak active energy of last 3 month
	4
	R
	

	003CH
	Flat active energy of last 3 month
	4
	R
	

	003EH
	Valley active energy of last 3 month
	4
	R
	

	0040H~0044H
	reserve

	0045H  
	status
	2
	R/W
	Bit0:0-E-no，1-EF-YES；                                              Bit1:0-forward，1-reverse;                                                                                           Bit3：0-PQS  1-RMS.


	0046H~0047H
	reserve

	0048H
	parity
	2
	R
	0000:None

0001:Odd
0002:Even

	0049H~004BH
	reserve

	004CH~004FH 
	Serial Number
	4
	R/W
	SN[14]
The last eight bits make up 0.

	0050H~0067H
	reserve

	0068H
	Current forward active total energy
	4
	R
	unit：0.01kWh

	006AH
	Current forward active spike energy
	4
	R
	

	006CH
	Current forward active peak energy
	4
	R
	

	006EH
	Current forward active flat energy
	4
	R
	

	0070H
	Current forward active valley energy
	4
	R
	

	0072H
	Current reversing active total energy
	4
	R
	

	0074H
	Current reversing active spike energy
	4
	R
	

	0076H
	Current reversing Active peak energy
	4
	R
	

	0078H
	Current reversing active flat energy
	4
	R
	

	007AH
	Current reversing Active valley energy
	4
	R
	

	007C~00AFH
	reserve

	00B0H
	Current total reactive energy
	4
	R
	unit：0.01kvarh

	00B2H
	Current spike reactive energy
	4
	R
	

	00B4H
	Current peak reactive energy
	4
	R
	

	00B6H
	Current flat reactive energy
	4
	R
	

	00B8H
	Current valley reactive energy
	4
	R
	

	00BAH
	Current forward reactive total energy
	4
	R
	

	00BCH
	Current forward reactive spike energy
	4
	R
	

	00BEH
	Current forward reactive peak energy
	4
	R
	

	00C0H
	Current forward reactive flat energy
	4
	R
	

	00C2H
	Current forward reactive valley energy
	4
	R
	

	00C4H
	Current reversing reactive total energy
	4
	R
	

	00C6H
	Current reversing reactive spike energy
	4
	R
	

	00C8H
	Current reversing reactive peak energy
	4
	R
	

	00CAH
	Current reversing reactive flat energy
	4
	R
	

	00CCH
	Current reversing reactive valley energy
	4
	R
	

	00CEH~52FFH
	reserve

	5300H 
	Voltage
	4
	R
	Float

	5302H 
	Current
	4
	R
	

	5304H 
	Active power
	4
	R
	

	5306H 
	Reactive power
	4
	R
	

	5308H 
	Apparent power
	4
	R
	

	530AH 
	power factor
	4
	R
	

	530CH
	Frequency
	4
	R
	

	530EH-

F008H
	reserve
	
	
	

	F009H
	Device model
	2
	R
	A200(HEX)

	0900H-091DH

	0900H
	L1 Phase Voltage
	4
	R
	Unit：V

	0902H
	L2 Phase Voltage
	4
	R
	V

	0904H
	L3 Phase Voltage
	4
	R
	V

	0906H
	L1 Current 
	4
	R
	A

	0908H
	L2 Current 
	4
	R
	A

	090AH
	L3 Current 
	4
	R
	A

	090CH
	Three-phase phase voltage average
	4
	R
	V

	090EH
	Three-phase phase current average
	4
	R
	A

	0910H
	Total Three-phase current
	4
	R
	A

	0912H
	Total active power
	4
	R
	W

	0914H
	L1-2 Line Voltage
	4
	R
	V

	0916H
	L2-3 Line Voltage
	4
	R
	V

	0918H
	L3-1 Line Voltage
	4
	R
	V

	091AH
	Three-phase line voltage average
	4
	R
	V

	091CH
	Total active power
	4
	R
	kWh

	091EH
	Maximum current value
	4
	R
	A

	0920H
	Maximum voltage value
	4
	R
	V


Headquarters: ZHEJIANG ETEK ELECTRICAL TECHNOLOGY CO.,LTD. 
Address: No. 288, Weishiqi Road, Economic Development Zone, Yueqing City, Zhejiang Province  
TEL.: 0086-0577-61713515 
Web-site: https://www.etek-electric.com/

_1234567893

_1234567895

_1234567896

_1234567897

_1234567894

_1234567891

_1234567892

_1234567890

